## AUC_task_summary
Task | DG-TIN AUC | Best feature baseline | Best baseline AUC | AUC gain
Detection | 0.903 | XGBoost | 0.879 | 0.024
Size 7-class | 0.957 | RF | 0.94 | 0.017
Size 3-bin | 0.947 | RF | 0.924 | 0.024
Depth 3-class | 0.792 | XGBoost | 0.762 | 0.03
Deep vs rest | 0.746 | RBF-SVM | 0.736 | 0.01

## AUC_full_model_comparison
Task | DG-TIN | LR | Peak amplitude | RBF-SVM | RF | XGBoost
Detection | 0.903 | 0.852 | 0.811 | 0.874 | 0.874 | 0.879
Size 7-class | 0.957 | 0.914 |  | 0.935 | 0.94 | 0.936
Size 3-bin | 0.947 | 0.882 |  | 0.913 | 0.924 | 0.92
Depth 3-class | 0.792 | 0.58 |  | 0.746 | 0.758 | 0.762
Deep vs rest | 0.746 | 0.554 |  | 0.736 | 0.718 | 0.711

## Readout_summary
Readout | Evaluation set | Primary metric | Clinical tolerance | Counts
Detection | All windows (n=3749) | Sensitivity 92.6% | Specificity 72.9% | TP 1282 / FP 641 / FN 103
Size | Detected positive windows (n=1282) | Exact 75.9% | Adj./exact 88.2%; Top-2 90.6% | Exact 973 / Adj. 158 / Severe 151
Depth | Detected positive windows (n=1282) | Exact 61.1% | Adj./exact 86.2%; Top-2 85.2% | Exact 783 / Adj. 322 / Far 177

## Size_7class
True size | Gate | Exact | Adj./exact | Top-2 | MAE cm | Under/Over | Most common pred.
0.25 cm | 186/211 | 98.4% | 98.4% | 98.4% | 0.02 | 0/3 | 0.25 cm (183)
0.5 cm | 334/355 | 89.8% | 90.7% | 98.5% | 0.07 | 0/34 | 0.5 cm (300)
0.75 cm | 127/155 | 78.0% | 83.5% | 84.3% | 0.16 | 0/28 | 0.75 cm (99)
1 cm | 165/171 | 64.8% | 92.1% | 85.5% | 0.1 | 30/28 | 1 cm (107)
1.25 cm | 165/168 | 36.4% | 70.3% | 67.3% | 0.2 | 34/71 | 1.25 cm (60)
1.5 cm | 175/192 | 68.0% | 82.9% | 94.3% | 0.17 | 30/26 | 1.5 cm (119)
1.75 cm | 130/133 | 80.8% | 96.9% | 96.2% | 0.05 | 25/0 | 1.75 cm (105)

## Depth_3class
True depth | Gate | Exact | Adj./exact | Top-2 | Under/Over | Most common pred.
Shallow 0.5-1.0 cm | 465/499 | 63.9% | 79.8% | 97.2% | 0/168 | shallow (297)
Middle 1.5-2.0 cm | 472/518 | 65.0% | 100.0% | 87.3% | 115/50 | middle (307)
Deep 2.5-3.0 cm | 345/368 | 51.9% | 75.9% | 66.1% | 166/0 | deep (179)

## Failure_mechanism_summary
Failure mode | Held-out evidence | Physical explanation | Clinical use
Detection false positive | 641 FPs; median contrast 20.4 vs 0.2 in TN, and P95 77.2 vs 3.0. | Bronchus, tissue edge or heterogeneous contact can create a local peak. | CT-prior plus repeat palpation; use size/depth plausibility.
Detection false negative | 103 FNs; median contrast 5.3 vs 24.5 in TP, with drift 0.044 vs 0.017. | Tangential touch, attenuation and local slip suppress center-border contrast. | Resweep CT-marked region; do not exclude on one low-P frame.
Size adjacent-bin error | Overall size Top-2 = 90.6%; the hardest exact bin is 1.25 cm (36.4% exact, 67.3% Top-2). | Larger, around-1-cm nodules deform under pressure, so amplitude and footprint do not scale perfectly. | Report an adjacent interval or Top-2 range instead of a single exact size class.
Depth under-call | Deep exact = 51.9%, deep Top-2 = 66.1%; overall depth Top-2 = 85.2%. | Deep responses are attenuated and can shift when the nodule slides inside the phantom/tissue. | Use depth as auxiliary guidance, not as a hard exclusion boundary.